# ISO 15189:2022-aligned method verification plan template

Clinical microbiology | All Things Microbiology

Template ID: ATM-MV-001 | Template version: 1.0 | Issued: 3 October 2026

Laboratory adaptation status: Draft - requires local review and approval.

# How to use this template

This original working template is aligned to ISO 15189:2022 for medical laboratories. It does not reproduce the standard, confer ISO approval or establish accreditation. Check your controlled copy of the standard and applicable regulator or accreditation-body guidance. Non-medical, food or environmental testing may require a different framework, such as ISO/IEC 17025 and applicable ISO 16140 guidance; assess that scope separately.

Type in the blank cells in the Word version. Print the PDF for a paper working copy. Duplicate study-design, acceptance-criteria and result tables for every applicable characteristic, method and specimen or organism group. Retain enough blank rows for the complete record and link to controlled raw data where a separate study workbook is used.

Complete and approve the plan before study testing. Define the intended use, clinical risks, study size and method-specific acceptance criteria locally. Sample counts, acceptance limits, allowed errors and confidence requirements are deliberately blank; there is no universal sample number or pass threshold in this template.

Use this plan for an unmodified, established method within its intended scope. Laboratory-developed, modified or out-of-scope methods require an appropriate validation programme. Select applicable studies and record a reason for each characteristic marked not applicable.

Use coded specimens and isolates in study records. Keep patient identifiers in the laboratory's approved records system. Completing the template does not authorise clinical use: an authorised person must review the evidence and approve the supported scope.

# 1. Document control and responsibilities

Complete the laboratory entries and maintain this plan as a controlled record.

| Field | Laboratory entry |
| --- | --- |
| Laboratory / department / site |  |
| Document ID / version / effective date |  |
| Method / examination name |  |
| Manufacturer / platform / software version / IFU revision |  |
| SOP ID / version |  |
| Planned study dates |  |
| Study lead / technical reviewer / authorising person |  |
| Raw-data location / retention period |  |
| Change history / superseded plan |  |

Approval table:

| Role | Name | Approval / signature | Date | Comments / conditions |
| --- | --- | --- | --- | --- |
| Prepared by |  |  |  |  |
| Technical / quality review |  |  |  |  |
| Plan authorised before testing |  |  |  |  |

# 2. Intended use, scope, and verification-or-validation decision

| Scope field | Laboratory entry |
| --- | --- |
| Clinical purpose and decisions supported |  |
| Method type: culture / identification / AST / qualitative NAAT / quantitative examination / other |  |
| Analytes / organisms / organism groups / antimicrobial agents |  |
| Specimen types / matrices / transport and storage conditions |  |
| Patient population / setting / inclusion and exclusion scope |  |
| Reported result, units, interpretation / clinical decision limits |  |
| Claimed range / detection capability / relevant manufacturer claims |  |
| Local operators, equipment, environment, workflow |  |
| Scope exclusions and limitations |  |

Record the scope decision before approving the study.

| Decision question | Yes / No / N/A | Evidence and implications |
| --- | --- | --- |
| Is this an established, previously validated method used without modification? |  |  |
| Are all local specimens, populations, intended uses and conditions within its established scope? |  |  |
| Are documented performance claims available and suitable for the laboratory’s intended use? |  |  |
| Is the method laboratory-developed, modified, or being used outside its established scope? |  |  |
| Are any unresolved scope or performance questions present? |  |  |

Route selected: [ ] Verification of an unmodified established method within its scope [ ] Validation required [ ] Technical review needed before proceeding.

Decision rationale / authoriser / date: ____________________________________________

For laboratory-developed, modified, or out-of-scope methods, complete an appropriate validation plan before clinical implementation. This verification template does not establish the full validation programme. Do not treat a change of specimen matrix, extraction method, algorithm, threshold, or intended use as automatically covered by a manufacturer’s original validation.

# 3. Prerequisites and clinical risk review

| Prerequisite | Evidence reference | Ready / Not ready / N/A | Gap, owner and due date |
| --- | --- | --- | --- |
| Controlled instructions and method SOP |  |  |  |
| Equipment qualification, maintenance and relevant calibration |  |  |  |
| Staff training / competency / authorisation |  |  |  |
| Reagent, medium, control and reference-material suitability |  |  |  |
| Comparator suitability and independent evidence |  |  |  |
| Biosafety, contamination prevention and waste procedures |  |  |  |
| Specimen suitability, consent / ethical arrangements where applicable |  |  |  |
| Result calculation, reporting and LIMS interface checks |  |  |  |
| QC and validity rules for study runs |  |  |  |

Risk table:

| Failure mode / affected result | Potential clinical harm | Likelihood / impact and rationale | Study element / control that addresses the risk | Residual risk / reviewer |
| --- | --- | --- | --- | --- |

Risk prompts include missed pathogen, false-positive contamination, misidentification, incorrect susceptibility category, near-threshold classification, quantitative bias, invalid results, delays and transcription errors. Use the laboratory's documented risk scale and rationale.

# 4. Performance characteristics to verify

Do not force every characteristic onto every method. Require an applicable/not-applicable decision with a rationale.

| Characteristic | Applicable / N/A and rationale | Source of claim or requirement | Planned study ID |
| --- | --- | --- | --- |
| Agreement with suitable comparator / reference |  |  |  |
| Repeatability / reproducibility appropriate to the result type |  |  |  |
| Detection / recovery capability or relevant analytical sensitivity |  |  |  |
| Specificity / selectivity / cross-reactivity where relevant |  |  |  |
| Reportable range / linearity / quantitation where relevant |  |  |  |
| Bias / trueness where relevant and reference suitable |  |  |  |
| Interference / carry-over / contamination where relevant |  |  |  |
| Specimen stability / handling conditions where relevant |  |  |  |
| Invalid, indeterminate, no-identification or no-result handling |  |  |  |
| Turnaround / reporting / interpretive rules where relevant |  |  |  |
| Measurement uncertainty evaluation or relevant information, where applicable |  |  |  |

For a qualitative result, assess repeatability or result consistency as appropriate; do not require a numerical CV. For a quantitative method, plan suitable precision/bias/range studies and assess measurement uncertainty where applicable. A laboratory should explain any claimed characteristic that it will not test locally and identify the evidence supporting that decision.

# 5. Study design and justified sample sizes

Repeat the following table for each study ID. Keep the study design separate from its results.

| Design field | Laboratory entry |
| --- | --- |
| Study ID / characteristic / objective |  |
| Intended clinical claim being verified |  |
| Specimen / isolate / reference material sources |  |
| Independent samples / isolates planned, by required category or stratum |  |
| Replicates / runs / days / operators / instruments / reagent lots |  |
| Challenge groups: positives, negatives, low-level or near-cut-off material, relevant organisms / resistance phenotypes |  |
| Inclusion / exclusion criteria and handling of unsuitable samples |  |
| Comparator / reference method, qualification and limitations |  |
| Pairing, masking, order of testing and independent interpretation |  |
| Handling / storage / freeze-thaw conditions |  |
| Statistical approach, denominator, confidence interval method if used |  |
| Sample size rationale: clinical risk, desired certainty, performance claim, available guidance and feasibility |  |
| Limits of the design / additional evidence or work needed |  |
| Data form / raw record references |  |

A specimen-stratum table is useful:

| Group / matrix / organism or antimicrobial | Independent samples planned | Replicates per sample | Selection rationale / source | Clinical risk addressed |
| --- | --- | --- | --- | --- |

The study must distinguish independent patient samples / isolates from technical replicates. Repeated testing of one isolate does not create additional independent clinical samples. Include clinically important uncommon groups and near-decision-limit material when relevant. A convenience panel alone may not support the proposed scope; record any coverage limits. Control material may support precision or detection studies but does not automatically establish patient-sample agreement.

# 6. Pre-agreed acceptance criteria and decision rules

This table is completed and approved before testing. Values are deliberately blank.

| Study ID / characteristic / group | Metric and exact denominator | Acceptance limit / allowed error | Source / rationale | Confidence or uncertainty requirement, if applicable | Decision / escalation rule |
| --- | --- | --- | --- | --- | --- |

| Rule to define before testing | Laboratory entry |
| --- | --- |
| Run validity / QC failure rule |  |
| Invalid / indeterminate / no-result handling and reporting |  |
| Repeats and allowed reason for repeat testing |  |
| Discordance investigation / independent resolution process |  |
| Predetermined exclusion rule and excluded-result reporting |  |
| Criteria for an inconclusive study / insufficient evidence |  |
| Handling of failed criteria and required additional work |  |
| Control of plan amendments: reason, approval, date and impact |  |

Retain original results when repeats or discrepancy investigations occur. Report original paired agreement separately from resolved findings; do not replace an unfavourable pair with a resolved result or remove it without a documented pre-agreed reason. Changing criteria after seeing results requires a controlled amendment and transparent justification, rather than a retrospective pass threshold.

# 7. Method-specific planning prompts

Use only the applicable row(s); these are prompts, not mandatory numerical thresholds.

| Method | What to specify in the plan |
| --- | --- |
| Culture / organism recovery | Specimen matrices, media and atmosphere, incubation conditions, target and non-target organisms, mixed growth where relevant, recovery/identity endpoints, clinically important low-level challenges, comparator limitations, and interpretation/reporting rules. |
| Identification | Claimed taxonomic/reporting level, representative species or groups, difficult/near-neighbour organisms, database/software version, correct identification versus misidentification versus no-identification, and an appropriate independent reference for discordance. |
| AST | Organism–antimicrobial combinations, reference method, local specimen/organism scope, applicable CLSI or EUCAST document and edition, breakpoint/version and category definitions, QC rules, quantitative MIC or zone results where relevant, categorical agreement and relevant error definitions/denominators. |
| Qualitative NAAT | Target(s), specimen matrices, extraction and internal controls, positive/negative agreement with a suitable comparator, low-level detection near the claimed limit, repeatability, inhibitors, cross-reactivity and contamination/carry-over where relevant, invalid/indeterminate outcomes, and cut-off interpretation. |
| Quantitative examination | Measurand and units, traceability/reference material, relevant range and decision levels, repeatability/intermediate precision, bias or suitable comparison, range/linearity and quantitation claims where applicable, interference/carry-over, uncertainty and reporting. |

AST details to include only if relevant:

| AST metric / rule | Exact local definition and denominator / approved criterion |
| --- | --- |
| Essential agreement for MIC results |  |
| Categorical agreement |  |
| Very major error / false-susceptible result: numerator and reference-resistant denominator |  |
| Major error / false-resistant result: numerator and reference-susceptible denominator |  |
| Other category discrepancies and applicable treatment of intermediate / increased-exposure categories |  |
| Treatment of off-scale MICs, technical uncertainty and unreportable combinations |  |
| Organism–drug-specific and overall summaries, including sparse denominators |  |

AST is not one homogeneous pooled claim. Report organism–drug groups and clinically important resistance categories where appropriate. A percentage with few reference-resistant isolates can conceal serious uncertainty; show the counts. CLSI and EUCAST category definitions cannot be mixed. EUCAST “I” means susceptible, increased exposure; it should not be casually relabelled as a generic intermediate category. State the chosen version because breakpoints change. Do not apply essential agreement to disc zone diameters as if they were MIC dilution results.

For NAAT and other qualitative comparisons, use positive / negative percentage agreement when the comparator is not a sufficiently established clinical reference. Do not rename comparator agreement as diagnostic sensitivity or specificity without a suitable clinical reference standard and study design. Cycle-threshold precision alone does not establish qualitative detection performance.

# 8. Execution record, deviations, and analysis

Raw-data record (repeat rows as needed):

| Coded sample / isolate ID and group | Date / run / operator | Lot / instrument | Candidate original result | Comparator original result | QC / validity | Repeat / resolution reference | Deviation / comment |
| --- | --- | --- | --- | --- | --- | --- | --- |

Separate disagreement/deviation record:

| Record ID / original finding | Investigation / evidence | Resolution and clinical significance | Corrective action / owner / due date | Closure / reviewer |
| --- | --- | --- | --- | --- |

Results summary:

| Study ID / characteristic / group | Planned / tested / excluded counts and reasons | Estimate / count / uncertainty or CI if specified | Approved criterion | Met / Not met / Inconclusive | Evidence reference / reviewer |
| --- | --- | --- | --- | --- | --- |

Show invalid, indeterminate, no-identification and no-result counts separately with their denominators. Do not silently remove these from an agreement analysis. Keep traceable coded IDs; do not put identifiable patient information in publicly shared template copies.

# 9. Technical conclusion and authorisation for intended use

| Review field | Laboratory entry |
| --- | --- |
| Verified claims and supported scope |  |
| Failed / inconclusive claims, limitations and unresolved questions |  |
| Deviations and their effect on interpretation |  |
| Additional study / corrective action required |  |
| Clinical risk review and any scope restrictions |  |
| SOP / reporting / training / QC changes needed |  |
| Overall decision: verified within stated scope / not verified / inconclusive |  |
| Authorised clinical scope and exclusions, if authorised |  |
| Implementation date / responsible person |  |

Review and release table:

| Role | Name | Decision / signature | Date | Conditions / evidence references |
| --- | --- | --- | --- | --- |
| Technical review |  |  |  |  |
| Quality review |  |  |  |  |
| Authorisation for stated intended use |  |  |  |  |

A completed form is not itself clinical authorisation. The authorised person must review the evidence and decide whether the defined scope is supported. Failed or unresolved criteria cannot be hidden behind a generic “pass”.

# 10. Implementation, ongoing assurance and change review

| Implementation / monitoring item | Requirement / evidence | Owner | Due / review date | Status |
| --- | --- | --- | --- | --- |
| Controlled SOP and interpretation / reporting rules |  |  |  |  |
| Staff competency and authorisation |  |  |  |  |
| QC plan and action rules |  |  |  |  |
| EQA / interlaboratory comparison or suitable alternative, where applicable |  |  |  |  |
| LIMS result mapping, calculations, flags and audit trail |  |  |  |  |
| Reagent / medium / control lot acceptance arrangements |  |  |  |  |
| Review of invalids, discordance and clinically important errors |  |  |  |  |
| Planned review and triggers for re-verification / validation assessment |  |  |  |  |

Change triggers should include changes in method, intended use, specimen type, instrument, critical software/database/algorithm, significant reagent changes, clinical decision limits or AST breakpoints, relocation/conditions, and persistent QC/EQA or performance concerns. Assess the effect of a change and document the needed work; not every lot change automatically requires a complete new method verification.

# 11. References and local guidance

Keep the laboratory's applicable editions controlled rather than assuming a universal numeric acceptance criterion.

| Reference / controlled edition | Relevant use / location |
| --- | --- |
| ISO 15189:2022 — Medical laboratories — Requirements for quality and competence | Primary alignment; verification, validation and validity of results. Official catalogue: https://www.iso.org/standard/76677.html |
| Manufacturer IFU / performance claims / revision |  |
| Laboratory SOP / verification policy / risk policy |  |
| Applicable national regulator / accreditation body guidance |  |
| Applicable CLSI / EUCAST method, QC and breakpoint documents and editions |  |
| ISO 20776 series, if applicable to AST study design |  |
| Applicable method-specific verification / statistics guidance |  |